## Slide 1
Assessment of Analytical Quality through Sigma metrics & its application for selection of
Westgard rule

## Slide 2: Objectives
What is six sigma?
How to estimate Sigma metrics of various biochemical analytes
Selection and application of Westgard rule based on sigma value.
Limitation of six sigma approch.

## Slide 3: What is six sigma?
Sigma Level | Accuracy | Long-Term ppm* Defects
1 | 30.85% | 691,462
2 | 69.1% | 308,538
3 | 99.33% | 66,807
4 | 99.38% | 6,210
5 | 99.977% | 233
6 | 99.99966% | 3.4
Sigma level and ppm defects or errors per Million Opportunities (DPMO)

## Slide 4: What is six sigma?
Six Sigma is a statistical approach.
Used to evaluate quality control performance
<2 (Unacceptable)
2-3 (Poor)
3-4 (Acceptable)
4-5 (Good)
5-6 (Excellent)
>6 (world class performance)

## Slide 5: Methodology
1. Both levels of IQC were analyzed each day. By using IQC data; Mean & S.D. were determined from which CV% was calculated for each month.
CV% = SD x 100 / Lab Mean
2. EQAS sample was run once in every month.
Bias (%) = (Mean of all laboratories using same instrument & method – Our Laboratory Mean) / (Mean of all laboratories using same instrument & method) x 100
3. The total permissible error (TEa) values were derived from CLIA-88 (Clinical Laboratories Improvement Act) recommendations.
4. Calculation of sigma metrics:
Sigma metrics (σ) = (TEa – Bias) / CV

## Slide 6: Table 1:Average Bias, Average CV% & sigma metrics calculated for march 2023
Parameter | TEa% | Average Bias% | Average CV% (L1) | σ (L1) | Average CV% (L2) | σ (L2)
Glucose | 10 | 1.4 | 2.61 | 3.295 | 2.67 | 3.221
Urea | 9 | 0.61 | 4.02 | 2.087 | 3.58 | 2.344
Creatinine | 15 | 0.2 | 2.33 | 6.352 | 2.33 | 6.352
TB | 20 | 0.6 | 6.11 | 3.175 | 2.57 | 7.549
DB | 20 | 3.4 | 7.89 | 2.104 | 3.39 | 4.897
SGPT | 20 | 4.2 | 4.01 | 3.94 | 2.40 | 6.583
SGOT | 20 | 1.9 | 2.43 | 7.449 | 2.37 | 7.637
ALP | 30 | 3.8 | 7.76 | 3.376 | 6.85 | 3.825
Total Protein | 10 | 1.05 | 3.73 | 2.399 | 3.11 | 2.878
Albumin | 10 | 1.3 | 3.84 | 2.266 | 3.79 | 2.296
Cholesterol | 10 | 1.7 | 2.62 | 3.168 | 3.43 | 2.42
Triglyceride | 25 | 1.9 | 3.83 | 6.031 | 4.18 | 5.526
HDL | 30 | 3.0 | 3.27 | 8.257 | 3.89 | 6.941
LDL | 20 | 6.6 | 2.57 | 5.214 | 3.54 | 3.785

## Slide 7: Table 1:Average Bias, Average CV% & sigma metrics calculated for may 2023
Parameter | TEa% | Average Bias% | Average CV% (L1) | σ (L1) | Average CV% (L2) | σ (L2)
Glucose | 10 | 2.7 | 1.37 | 5.33 | 0.67 | 10.90
Urea | 9 | 3.1 | 6.18 | 0.95 | 8.78 | 0.67
Creatinine | 15 | 15.4 | 11.06 | -0.04 | 10.99 | -0.04
TB | 20 | 13.4 | 8.83 | 0.75 | 4.53 | 1.46
DB | 20 | 3.9 | 4.69 | 3.43 | 4.97 | 3.24
SGPT | 20 | 7 | 3.35 | 3.88 | 6.5 | 2.00
SGOT | 20 | 3.9 | 1.71 | 9.42 | 5.7 | 2.82
ALP | 30 | 10.9 | 19.23 | 0.99 | 17.06 | 1.12
Total Protein | 10 | 7.2 | 3.7 | 0.76 | 4.3 | 0.65
Albumin | 10 | 1.5 | 2.63 | 3.23 | 1.96 | 4.34
Cholesterol | 10 | 2.3 | 2.92 | 2.64 | 2.53 | 3.04
Triglyceride | 25 | 0.9 | 6.27 | 3.84 | 3.07 | 7.85
HDL | 30 | 0.6 | 4.13 | 7.12 | 4.07 | 7.22
LDL | 20 | 9.9 | 3.41 | 2.96 | 2.24 | 4.51

## Slide 8: Selection of Westgard rule based on Sigma value
>6σ
Evaluate with 1 QC/day that is alternating levels between days and follow 1-3 s Westgard rule.
4 σ - 6 σ
Evaluate with two levels of QC /day and follow 1‑3 s, 2‑2 s, R4 s Westgard multirules.
3 σ - 4 σ
Evaluate with two levels of QC /day and follow 1‑3 s, 2‑2 s, R4s, and 4‑1 s Westgard multirules.
< 3 σ
Evaluate with 3 levels of QC 3 times a day.
Root cause analysis should be performed.

## Slide 9: Limitations of sigma approach
CV% is derived after exclusion of QC violating 13s or more.

## Slide 10: Assuring Good Laboratory Practices
Thank you